## Sheet1
Gretag i1pro driver vs. ArgyllCMS V1.2.0 |  |  |  | DE is CIE 1976
| Overall: |  | Avg | Max
| DE spec Lab: |  | 0.0057669 | 0.011687 | Difference between Gretag & Argyll spectral to XYZ/Lab
| DE to Gretag XYZ.: |  | 0.026483 | 0.048006 | Difference between Argyll Measurement and Gretag Lab
| DE to Gretag spec.: |  | 0.028021 | 0.050861 | Difference between Argyll Measurement and Gretag measurenent from spectral
| DE to Argyll Norm: |  | 0.158 | 0.35346 | Difference between Argyll normal and Hi-res measurement
Cal tile
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 85.8017 | 89.23 | 70.7037 | 95.6765 | -0.43651 | 2.5645
Gretag spec. | DE spec Lab: | 0.0083947 |  | 85.8039 | 89.23 | 70.6955 | 95.6765 | -0.43251 | 2.5719 | 68.204 | 73.581 | 77.688 | 80.477 | 82.262 | 83.63 | 84.494 | 85.253 | 86.279 | 86.947 | 87.286 | 87.796 | 88.204 | 88.448 | 88.84 | 89.082 | 89.273 | 89.508 | 89.348 | 89.677 | 89.454 | 89.435 | 89.677 | 89.742 | 89.718 | 89.728 | 89.817 | 89.478 | 89.673 | 90.192 | 90.242 | 90.084 | 90.379 | 90.603 | 90.559 | 90.682
| DE to Gretag XYZ.: | 0.008765
Argyll | DE to Gretag spec.: | 0.011935 |  | 85.7855 | 89.2112 | 70.6904 | 95.6686 | -0.43296 | 2.563 | 68.096 | 73.643 | 77.603 | 80.429 | 82.216 | 83.638 | 84.494 | 85.265 | 86.266 | 86.923 | 87.291 | 87.779 | 88.217 | 88.432 | 88.832 | 89.074 | 89.245 | 89.469 | 89.328 | 89.644 | 89.452 | 89.429 | 89.646 | 89.732 | 89.698 | 89.694 | 89.779 | 89.455 | 89.644 | 90.159 | 90.209 | 90.058 | 90.349 | 90.564 | 90.509 | 90.631
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.010046 |  | 85.7893 | 89.2149 | 70.7045 | 95.6702 | -0.4327 | 2.553 | 66.262 | 68.223 | 70.122 | 71.968 | 73.639 | 75.148 | 76.5 | 77.676 | 78.727 | 79.676 | 80.479 | 81.14 | 81.735 | 82.262 | 82.766 | 83.248 | 83.653 | 83.988 | 84.273 | 84.522 | 84.768 | 85.023 | 85.292 | 85.621 | 85.963 | 86.283 | 86.525 | 86.752 | 86.925 | 87.065 | 87.18 | 87.317 | 87.458 | 87.626 | 87.796 | 87.947 | 88.088 | 88.217 | 88.301 | 88.365 | 88.445 | 88.566 | 88.7 | 88.836 | 88.943 | 89.016 | 89.087 | 89.146 | 89.192 | 89.25 | 89.34 | 89.419 | 89.452 | 89.417 | 89.367 | 89.352 | 89.435 | 89.553 | 89.629 | 89.611 | 89.543 | 89.46 | 89.421 | 89.409 | 89.433 | 89.498 | 89.571 | 89.64 | 89.688 | 89.719 | 89.74 | 89.723 | 89.713 | 89.697 | 89.688 | 89.693 | 89.701 | 89.739 | 89.773 | 89.762 | 89.672 | 89.559 | 89.465 | 89.455 | 89.519 | 89.645 | 89.818 | 89.986 | 90.139 | 90.228 | 90.235 | 90.201 | 90.142 | 90.089 | 90.082 | 90.139 | 90.227 | 90.338 | 90.424 | 90.501 | 90.552 | 90.551 | 90.52 | 90.506 | 90.532 | 90.565 | 90.628 | 90.679 | 90.741 | 90.816 | 90.869 | 90.896 | 90.92
Dark Skin
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 11.5988 | 10.0632 | 4.8847 | 37.9557 | 14.2551 | 15.0724
Gretag spec. | DE spec Lab: | 0.0035186 |  | 11.5994 | 10.0634 | 4.8844 | 37.9561 | 14.2577 | 15.0747 | 5.107 | 5.718 | 6.663 | 7.142 | 6.883 | 6.362 | 5.927 | 5.65 | 5.566 | 5.576 | 5.721 | 6.019 | 6.388 | 6.782 | 7.32 | 7.739 | 8.016 | 8.296 | 8.831 | 9.898 | 11.286 | 12.588 | 13.569 | 14.338 | 15.076 | 15.886 | 16.929 | 18.268 | 20.054 | 22.024 | 24.04 | 26.127 | 28.473 | 31.113 | 34.268 | 38.294
| DE to Gretag XYZ.: | 0.027429
Argyll | DE to Gretag spec.: | 0.027231 |  | 11.6114 | 10.0729 | 4.8931 | 37.9729 | 14.2705 | 15.0576 | 5.195 | 5.75 | 6.656 | 7.082 | 6.848 | 6.354 | 5.945 | 5.666 | 5.572 | 5.607 | 5.741 | 6.029 | 6.397 | 6.799 | 7.319 | 7.742 | 8.021 | 8.302 | 8.83 | 9.906 | 11.303 | 12.605 | 13.586 | 14.357 | 15.092 | 15.899 | 16.931 | 18.287 | 20.074 | 22.047 | 24.066 | 26.149 | 28.5 | 31.136 | 34.289 | 38.345
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.061698 |  | 11.6042 | 10.0591 | 4.8877 | 37.9482 | 14.3254 | 15.0438 | 4.664 | 4.933 | 5.138 | 5.33 | 5.649 | 5.959 | 6.344 | 6.682 | 6.932 | 7.091 | 7.172 | 7.166 | 7.063 | 6.937 | 6.774 | 6.562 | 6.383 | 6.195 | 6.036 | 5.923 | 5.827 | 5.706 | 5.655 | 5.597 | 5.582 | 5.576 | 5.558 | 5.575 | 5.578 | 5.605 | 5.663 | 5.718 | 5.809 | 5.894 | 6.023 | 6.147 | 6.267 | 6.381 | 6.514 | 6.629 | 6.767 | 6.92 | 7.119 | 7.311 | 7.462 | 7.609 | 7.739 | 7.836 | 7.935 | 8.014 | 8.088 | 8.165 | 8.264 | 8.398 | 8.57 | 8.784 | 9.076 | 9.463 | 9.876 | 10.326 | 10.813 | 11.311 | 11.789 | 12.23 | 12.637 | 12.98 | 13.304 | 13.591 | 13.854 | 14.1 | 14.342 | 14.591 | 14.826 | 15.074 | 15.332 | 15.597 | 15.877 | 16.181 | 16.533 | 16.918 | 17.31 | 17.759 | 18.276 | 18.896 | 19.501 | 20.086 | 20.708 | 21.344 | 22.024 | 22.737 | 23.416 | 24.078 | 24.713 | 25.387 | 26.121 | 26.898 | 27.687 | 28.492 | 29.3 | 30.147 | 31.074 | 32.072 | 33.144 | 34.287 | 35.527 | 36.914 | 38.383 | 39.869 | 41.332 | 42.807 | 44.204 | 45.305 | 45.923
Light Skin
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 39.2838 | 35.5087 | 20.0292 | 66.1428 | 16.6043 | 16.8534
Gretag spec. | DE spec Lab: | 0.0063944 |  | 39.2851 | 35.5093 | 20.0267 | 66.1433 | 16.6067 | 16.8594 | 10.573 | 14.08 | 18.487 | 20.861 | 21.271 | 21.547 | 22.021 | 22.819 | 24.143 | 25.814 | 27.364 | 28.358 | 29.273 | 31.189 | 31.973 | 30.131 | 29.32 | 29.751 | 29.415 | 30.078 | 34.352 | 41.224 | 47.251 | 51.045 | 53.237 | 54.689 | 56.123 | 57.683 | 59.7 | 62 | 64.802 | 67.92 | 71.355 | 74.501 | 77.082 | 79.164
| DE to Gretag XYZ.: | 0.026487
Argyll | DE to Gretag spec.: | 0.031796 |  | 39.2931 | 35.5229 | 20.0462 | 66.1538 | 16.5865 | 16.8372 | 10.508 | 14.177 | 18.547 | 20.881 | 21.317 | 21.556 | 21.991 | 22.841 | 24.175 | 25.87 | 27.41 | 28.385 | 29.311 | 31.215 | 31.992 | 30.157 | 29.336 | 29.767 | 29.425 | 30.084 | 34.355 | 41.236 | 47.258 | 51.048 | 53.235 | 54.686 | 56.126 | 57.681 | 59.721 | 62.034 | 64.814 | 67.937 | 71.379 | 74.496 | 77.076 | 79.188
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.068917 |  | 39.2615 | 35.4873 | 19.9943 | 66.1263 | 16.6053 | 16.8975 | 9.471 | 10.039 | 10.825 | 12.21 | 13.764 | 15.503 | 17.132 | 18.585 | 19.752 | 20.52 | 20.913 | 21.122 | 21.234 | 21.294 | 21.361 | 21.417 | 21.522 | 21.618 | 21.734 | 21.923 | 22.168 | 22.441 | 22.786 | 23.176 | 23.559 | 24.06 | 24.587 | 25.147 | 25.757 | 26.381 | 26.928 | 27.419 | 27.822 | 28.118 | 28.393 | 28.582 | 28.804 | 29.147 | 29.696 | 30.439 | 31.281 | 32.019 | 32.406 | 32.304 | 31.712 | 30.859 | 30.057 | 29.505 | 29.224 | 29.195 | 29.361 | 29.646 | 29.863 | 29.885 | 29.687 | 29.344 | 29.129 | 29.228 | 29.729 | 30.691 | 32.176 | 34.121 | 36.381 | 38.82 | 41.298 | 43.602 | 45.639 | 47.376 | 48.864 | 50.1 | 51.13 | 51.984 | 52.691 | 53.269 | 53.777 | 54.246 | 54.681 | 55.126 | 55.623 | 56.112 | 56.549 | 57.028 | 57.669 | 58.484 | 59.14 | 59.712 | 60.352 | 61.1 | 61.931 | 62.858 | 63.831 | 64.807 | 65.774 | 66.8 | 67.956 | 69.135 | 70.281 | 71.415 | 72.473 | 73.508 | 74.507 | 75.469 | 76.363 | 77.184 | 77.907 | 78.587 | 79.246 | 79.819 | 80.338 | 80.776 | 81.123 | 81.368 | 81.105
Blue Sky
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 17.719 | 19.4902 | 26.9831 | 51.2558 | -5.6282 | -21.8452
Gretag spec. | DE spec Lab: | 0.0047636 |  | 17.7191 | 19.49 | 26.9803 | 51.2556 | -5.6265 | -21.8407 | 11.425 | 17.02 | 25.453 | 32.02 | 34.238 | 34.711 | 34.563 | 34.191 | 33.459 | 32.266 | 30.801 | 29.095 | 27.315 | 25.535 | 23.588 | 21.821 | 20.753 | 19.967 | 18.987 | 17.938 | 16.919 | 16.118 | 15.444 | 14.802 | 14.071 | 13.291 | 12.608 | 12.021 | 11.539 | 11.158 | 10.852 | 10.605 | 10.462 | 10.377 | 10.344 | 10.364
| DE to Gretag XYZ.: | 0.031347
Argyll | DE to Gretag spec.: | 0.033056 |  | 17.7364 | 19.5049 | 27.0102 | 51.2727 | -5.608 | -21.8622 | 11.469 | 17.098 | 25.52 | 32.092 | 34.327 | 34.747 | 34.621 | 34.238 | 33.48 | 32.286 | 30.823 | 29.118 | 27.345 | 25.544 | 23.587 | 21.829 | 20.774 | 19.985 | 19.002 | 17.951 | 16.94 | 16.137 | 15.456 | 14.82 | 14.078 | 13.307 | 12.627 | 12.016 | 11.542 | 11.153 | 10.852 | 10.614 | 10.477 | 10.382 | 10.355 | 10.384
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.042498 |  | 17.7347 | 19.5039 | 27.034 | 51.2716 | -5.6125 | -21.9044 | 9.973 | 10.76 | 11.94 | 13.841 | 16.103 | 19.062 | 22.358 | 25.558 | 28.379 | 30.692 | 32.38 | 33.38 | 34.002 | 34.41 | 34.589 | 34.688 | 34.745 | 34.745 | 34.699 | 34.591 | 34.507 | 34.405 | 34.264 | 34.101 | 33.824 | 33.535 | 33.198 | 32.796 | 32.347 | 31.905 | 31.401 | 30.898 | 30.341 | 29.71 | 29.118 | 28.521 | 27.907 | 27.324 | 26.759 | 26.186 | 25.582 | 24.938 | 24.273 | 23.641 | 22.966 | 22.309 | 21.785 | 21.389 | 21.052 | 20.767 | 20.519 | 20.258 | 19.986 | 19.703 | 19.387 | 19.014 | 18.628 | 18.279 | 17.93 | 17.554 | 17.208 | 16.921 | 16.655 | 16.39 | 16.128 | 15.901 | 15.662 | 15.466 | 15.254 | 15.045 | 14.834 | 14.598 | 14.338 | 14.069 | 13.816 | 13.556 | 13.287 | 13.038 | 12.808 | 12.603 | 12.38 | 12.161 | 11.985 | 11.865 | 11.708 | 11.529 | 11.382 | 11.237 | 11.122 | 11.022 | 10.943 | 10.847 | 10.742 | 10.665 | 10.6 | 10.541 | 10.513 | 10.474 | 10.428 | 10.383 | 10.369 | 10.367 | 10.366 | 10.351 | 10.358 | 10.352 | 10.383 | 10.39 | 10.406 | 10.422 | 10.456 | 10.515 | 10.514
Foliage
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 10.9997 | 13.3948 | 5.5539 | 43.3521 | -13.3301 | 20.9678
Gretag spec. | DE spec Lab: | 0.0047636 |  | 11.0001 | 13.3945 | 5.5532 | 43.3518 | -13.3257 | 20.971 | 4.509 | 4.958 | 5.202 | 5.354 | 5.47 | 5.622 | 5.865 | 6.268 | 6.666 | 6.995 | 7.255 | 7.552 | 8.272 | 10.632 | 14.931 | 17.888 | 18.031 | 16.463 | 14.459 | 13.214 | 12.821 | 12.104 | 11.11 | 10.401 | 10.32 | 10.495 | 10.369 | 10.099 | 10.042 | 10.56 | 12.377 | 16.451 | 22.557 | 28.545 | 32.368 | 34.216
| DE to Gretag XYZ.: | 0.026337
Argyll | DE to Gretag spec.: | 0.025952 |  | 11.0123 | 13.4097 | 5.5589 | 43.3742 | -13.3331 | 20.9818 | 4.869 | 5.108 | 5.277 | 5.409 | 5.443 | 5.606 | 5.891 | 6.274 | 6.677 | 7.005 | 7.254 | 7.55 | 8.275 | 10.656 | 14.953 | 17.899 | 18.045 | 16.479 | 14.483 | 13.232 | 12.836 | 12.121 | 11.129 | 10.415 | 10.332 | 10.499 | 10.363 | 10.093 | 10.048 | 10.56 | 12.38 | 16.481 | 22.587 | 28.597 | 32.386 | 34.23
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.077024 |  | 11.014 | 13.4215 | 5.5563 | 43.3915 | -13.3951 | 21.024 | 4.997 | 4.959 | 4.815 | 4.962 | 5 | 5.071 | 5.186 | 5.263 | 5.309 | 5.349 | 5.387 | 5.446 | 5.491 | 5.515 | 5.518 | 5.523 | 5.597 | 5.667 | 5.764 | 5.857 | 5.997 | 6.128 | 6.281 | 6.441 | 6.567 | 6.696 | 6.802 | 6.913 | 7.008 | 7.091 | 7.172 | 7.255 | 7.351 | 7.429 | 7.531 | 7.672 | 7.844 | 8.155 | 8.627 | 9.374 | 10.441 | 11.812 | 13.425 | 15.003 | 16.344 | 17.403 | 18.101 | 18.429 | 18.465 | 18.226 | 17.772 | 17.173 | 16.494 | 15.826 | 15.144 | 14.428 | 13.777 | 13.328 | 13.117 | 13.073 | 13.017 | 12.883 | 12.703 | 12.471 | 12.174 | 11.825 | 11.468 | 11.116 | 10.797 | 10.536 | 10.357 | 10.268 | 10.25 | 10.301 | 10.411 | 10.509 | 10.555 | 10.531 | 10.467 | 10.378 | 10.264 | 10.166 | 10.089 | 10.064 | 10.039 | 10.029 | 10.084 | 10.217 | 10.456 | 10.87 | 11.449 | 12.238 | 13.278 | 14.616 | 16.318 | 18.284 | 20.428 | 22.67 | 24.892 | 26.959 | 28.867 | 30.427 | 31.676 | 32.636 | 33.357 | 33.922 | 34.34 | 34.7 | 34.915 | 35.106 | 35.303 | 35.367 | 35.192
Blue Flower
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 25.175 | 24.0773 | 34.4706 | 56.1651 | 8.5167 | -25.1019
Gretag spec. | DE spec Lab: | 0.0052514 |  | 25.1754 | 24.0774 | 34.4672 | 56.1652 | 8.5183 | -25.0969 | 11.962 | 18.42 | 29.781 | 40.233 | 44.211 | 44.964 | 44.69 | 43.848 | 42.66 | 41.026 | 38.925 | 36.903 | 34.809 | 31.933 | 27.653 | 23.543 | 21.341 | 20.692 | 20.006 | 19.307 | 19.846 | 21.36 | 22.831 | 23.998 | 24.12 | 24.284 | 26.293 | 31.144 | 38.761 | 46.683 | 52.518 | 55.731 | 57.359 | 58.018 | 58.327 | 58.58
| DE to Gretag XYZ.: | 0.023261
Argyll | DE to Gretag spec.: | 0.026432 |  | 25.1945 | 24.0968 | 34.5026 | 56.1846 | 8.5153 | -25.1146 | 11.837 | 18.514 | 29.898 | 40.287 | 44.262 | 45.04 | 44.72 | 43.884 | 42.698 | 41.082 | 38.966 | 36.924 | 34.85 | 31.956 | 27.68 | 23.567 | 21.37 | 20.707 | 20.02 | 19.314 | 19.858 | 21.383 | 22.853 | 24.008 | 24.132 | 24.284 | 26.306 | 31.18 | 38.794 | 46.727 | 52.556 | 55.764 | 57.397 | 58.029 | 58.359 | 58.6
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.082848 |  | 25.1773 | 24.0809 | 34.5399 | 56.1687 | 8.5113 | -25.1958 | 10.138 | 11.149 | 12.535 | 14.36 | 17.195 | 20.823 | 25.19 | 29.798 | 34.173 | 37.912 | 40.719 | 42.518 | 43.709 | 44.469 | 44.801 | 44.963 | 45.031 | 45.038 | 44.909 | 44.728 | 44.51 | 44.226 | 43.929 | 43.572 | 43.162 | 42.769 | 42.311 | 41.75 | 41.211 | 40.56 | 39.804 | 39.008 | 38.242 | 37.539 | 36.93 | 36.255 | 35.581 | 34.865 | 34.077 | 33.192 | 32.131 | 30.818 | 29.31 | 27.772 | 26.236 | 24.725 | 23.481 | 22.542 | 21.805 | 21.251 | 20.904 | 20.735 | 20.691 | 20.658 | 20.461 | 20.05 | 19.608 | 19.309 | 19.175 | 19.196 | 19.401 | 19.781 | 20.275 | 20.827 | 21.395 | 21.931 | 22.398 | 22.83 | 23.268 | 23.729 | 24.12 | 24.32 | 24.291 | 24.117 | 23.963 | 23.964 | 24.148 | 24.548 | 25.176 | 26.09 | 27.327 | 28.955 | 31.038 | 33.541 | 36.175 | 38.862 | 41.625 | 44.344 | 46.928 | 49.248 | 51.21 | 52.773 | 54.042 | 55.061 | 55.898 | 56.594 | 57.093 | 57.453 | 57.724 | 57.866 | 57.987 | 58.159 | 58.304 | 58.407 | 58.436 | 58.459 | 58.578 | 58.71 | 58.762 | 58.846 | 58.921 | 58.945 | 58.616
Blueish Green
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 31.3171 | 42.6637 | 34.8047 | 71.3261 | -32.7093 | 0.55624
Gretag spec. | DE spec Lab: | 0.0064495 |  | 31.3173 | 42.663 | 34.8001 | 71.3256 | -32.7064 | 0.562 | 11.679 | 17.177 | 25.192 | 31.492 | 34.028 | 35.469 | 36.924 | 39.057 | 42.069 | 45.594 | 49.008 | 52.824 | 56.662 | 58.973 | 59.119 | 57.253 | 54.236 | 49.951 | 44.792 | 39.84 | 35.02 | 30.407 | 26.233 | 23.339 | 21.782 | 21.012 | 20.583 | 20.311 | 20.586 | 21.414 | 22.546 | 23.626 | 24.468 | 24.668 | 24.36 | 24.402
| DE to Gretag XYZ.: | 0.044928
Argyll | DE to Gretag spec.: | 0.048363 |  | 31.3424 | 42.687 | 34.8479 | 71.342 | -32.6855 | 0.52164 | 11.813 | 17.234 | 25.312 | 31.594 | 34.094 | 35.532 | 36.999 | 39.106 | 42.134 | 45.655 | 49.042 | 52.87 | 56.695 | 59.011 | 59.139 | 57.277 | 54.264 | 49.981 | 44.81 | 39.862 | 35.034 | 30.431 | 26.249 | 23.359 | 21.803 | 21.039 | 20.594 | 20.325 | 20.612 | 21.453 | 22.594 | 23.659 | 24.496 | 24.683 | 24.367 | 24.391
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.26949 |  | 31.2938 | 42.7011 | 34.7506 | 71.3516 | -32.9048 | 0.67792 | 10.558 | 11.131 | 12.265 | 13.978 | 16.274 | 19.121 | 22.238 | 25.322 | 28.042 | 30.211 | 31.778 | 32.791 | 33.523 | 34.087 | 34.515 | 34.97 | 35.452 | 35.9 | 36.38 | 36.901 | 37.468 | 38.143 | 39.015 | 39.896 | 40.823 | 41.911 | 43.081 | 44.303 | 45.497 | 46.591 | 47.72 | 48.922 | 50.126 | 51.438 | 52.822 | 54.184 | 55.541 | 56.798 | 57.844 | 58.651 | 59.196 | 59.457 | 59.497 | 59.331 | 58.875 | 58.163 | 57.404 | 56.581 | 55.587 | 54.436 | 53.108 | 51.626 | 50.003 | 48.323 | 46.576 | 44.774 | 43.007 | 41.371 | 39.783 | 38.166 | 36.541 | 34.999 | 33.494 | 31.973 | 30.424 | 28.907 | 27.473 | 26.155 | 24.993 | 24.021 | 23.249 | 22.631 | 22.131 | 21.724 | 21.428 | 21.201 | 21.003 | 20.838 | 20.706 | 20.563 | 20.425 | 20.297 | 20.264 | 20.356 | 20.434 | 20.551 | 20.766 | 21.051 | 21.387 | 21.809 | 22.209 | 22.593 | 22.944 | 23.314 | 23.676 | 24.011 | 24.309 | 24.544 | 24.68 | 24.727 | 24.703 | 24.623 | 24.503 | 24.328 | 24.213 | 24.159 | 24.262 | 24.54 | 24.912 | 25.443 | 26.105 | 26.777 | 27.327
Orange
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 39.8178 | 30.6876 | 4.7727 | 62.243 | 35.086 | 57.5473
Gretag spec. | DE spec Lab: | 0.0053396 |  | 39.8201 | 30.6889 | 4.7722 | 62.2441 | 35.0884 | 57.552 | 5.719 | 5.766 | 5.865 | 5.742 | 5.578 | 5.523 | 5.469 | 5.506 | 5.55 | 5.655 | 5.8 | 5.989 | 6.326 | 7.064 | 8.799 | 12.092 | 16.716 | 22.244 | 28.732 | 36.796 | 44.508 | 50.208 | 53.754 | 55.965 | 57.473 | 58.586 | 59.663 | 60.534 | 61.443 | 62.084 | 62.623 | 63.13 | 63.888 | 64.465 | 65.007 | 65.561
| DE to Gretag XYZ.: | 0.033542
Argyll | DE to Gretag spec.: | 0.036846 |  | 39.8295 | 30.7051 | 4.7795 | 62.2579 | 35.0581 | 57.5363 | 5.793 | 5.835 | 5.873 | 5.792 | 5.654 | 5.534 | 5.496 | 5.497 | 5.549 | 5.652 | 5.799 | 6.006 | 6.338 | 7.076 | 8.811 | 12.115 | 16.753 | 22.282 | 28.755 | 36.811 | 44.526 | 50.209 | 53.741 | 55.974 | 57.476 | 58.602 | 59.664 | 60.541 | 61.451 | 62.067 | 62.642 | 63.151 | 63.911 | 64.493 | 65.031 | 65.582
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.20096 |  | 39.8666 | 30.6833 | 4.7658 | 62.2394 | 35.2536 | 57.5787 | 5.907 | 5.545 | 5.534 | 5.714 | 5.707 | 5.712 | 5.776 | 5.755 | 5.82 | 5.823 | 5.785 | 5.733 | 5.623 | 5.604 | 5.601 | 5.543 | 5.509 | 5.486 | 5.502 | 5.477 | 5.497 | 5.495 | 5.486 | 5.508 | 5.517 | 5.522 | 5.573 | 5.608 | 5.624 | 5.675 | 5.738 | 5.808 | 5.86 | 5.914 | 5.983 | 6.068 | 6.169 | 6.284 | 6.465 | 6.705 | 6.996 | 7.391 | 7.941 | 8.627 | 9.476 | 10.547 | 11.838 | 13.307 | 14.912 | 16.63 | 18.415 | 20.255 | 22.199 | 24.21 | 26.327 | 28.662 | 31.282 | 34.076 | 36.888 | 39.646 | 42.297 | 44.75 | 46.912 | 48.794 | 50.399 | 51.734 | 52.856 | 53.818 | 54.624 | 55.372 | 56.038 | 56.59 | 57.075 | 57.508 | 57.902 | 58.277 | 58.639 | 58.981 | 59.335 | 59.684 | 59.956 | 60.206 | 60.566 | 61.061 | 61.374 | 61.481 | 61.638 | 61.81 | 62.008 | 62.239 | 62.49 | 62.658 | 62.789 | 62.946 | 63.164 | 63.433 | 63.71 | 63.951 | 64.108 | 64.255 | 64.459 | 64.673 | 64.904 | 65.076 | 65.209 | 65.387 | 65.566 | 65.791 | 65.99 | 66.152 | 66.299 | 66.338 | 66.005
Purpleish Blue
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 12.3187 | 11.4858 | 28.2974 | 40.3869 | 8.7803 | -42.7867
Gretag spec. | DE spec Lab: | 0.0070233 |  | 12.3184 | 11.4856 | 28.293 | 40.3866 | 8.779 | -42.7798 | 10.185 | 14.858 | 22.233 | 28.969 | 32.587 | 35.039 | 37.254 | 38.779 | 37.865 | 34.986 | 30.942 | 25.815 | 21.005 | 17.255 | 14.043 | 11.739 | 10.635 | 9.886 | 9.002 | 8.367 | 8.264 | 8.541 | 8.73 | 8.617 | 8.465 | 8.685 | 9.48 | 10.685 | 11.962 | 12.572 | 12.25 | 11.661 | 11.545 | 11.856 | 12.477 | 13.424
| DE to Gretag XYZ.: | 0.025735
Argyll | DE to Gretag spec.: | 0.020487 |  | 12.3218 | 11.491 | 28.2951 | 40.3954 | 8.7647 | -42.7681 | 10.003 | 14.843 | 22.197 | 28.933 | 32.562 | 35.034 | 37.27 | 38.768 | 37.875 | 34.997 | 30.945 | 25.815 | 21.019 | 17.271 | 14.052 | 11.753 | 10.641 | 9.889 | 9.004 | 8.369 | 8.268 | 8.543 | 8.719 | 8.625 | 8.482 | 8.685 | 9.484 | 10.695 | 11.981 | 12.57 | 12.259 | 11.669 | 11.546 | 11.855 | 12.481 | 13.444
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.22262 |  | 12.3339 | 11.4848 | 28.3956 | 40.3852 | 8.8906 | -42.9514 | 8.629 | 9.411 | 10.575 | 12.08 | 14.038 | 16.513 | 19.362 | 22.224 | 24.849 | 27.168 | 29.147 | 30.612 | 31.688 | 32.538 | 33.258 | 34.097 | 34.972 | 35.74 | 36.488 | 37.279 | 38.1 | 38.707 | 39.007 | 38.994 | 38.671 | 38.141 | 37.307 | 36.323 | 35.227 | 34.048 | 32.715 | 31.21 | 29.533 | 27.653 | 25.79 | 24.034 | 22.369 | 20.886 | 19.568 | 18.388 | 17.274 | 16.162 | 15.074 | 14.052 | 13.124 | 12.302 | 11.688 | 11.254 | 10.912 | 10.624 | 10.384 | 10.147 | 9.892 | 9.628 | 9.324 | 8.99 | 8.687 | 8.463 | 8.322 | 8.232 | 8.202 | 8.249 | 8.328 | 8.436 | 8.539 | 8.653 | 8.722 | 8.748 | 8.731 | 8.697 | 8.643 | 8.569 | 8.503 | 8.449 | 8.455 | 8.518 | 8.644 | 8.843 | 9.118 | 9.455 | 9.832 | 10.248 | 10.715 | 11.223 | 11.671 | 12.058 | 12.355 | 12.556 | 12.642 | 12.621 | 12.488 | 12.296 | 12.035 | 11.816 | 11.631 | 11.53 | 11.504 | 11.508 | 11.562 | 11.673 | 11.818 | 12.021 | 12.233 | 12.477 | 12.724 | 13.035 | 13.396 | 13.839 | 14.42 | 15.102 | 15.896 | 16.711 | 17.428
Moderate Red
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 30.2693 | 20.0675 | 10.7263 | 51.9134 | 47.0893 | 15.7681
Gretag spec. | DE spec Lab: | 0.0044986 |  | 30.27 | 20.0679 | 10.7251 | 51.9139 | 47.0897 | 15.7726 | 9.88 | 12.356 | 14.446 | 14.899 | 14.274 | 13.78 | 13.525 | 13.331 | 13.108 | 12.883 | 12.347 | 11.436 | 10.742 | 10.341 | 9.748 | 9.25 | 9.373 | 10.089 | 10.844 | 11.58 | 15.464 | 25.85 | 39.409 | 49.986 | 55.703 | 58.135 | 59.133 | 59.424 | 59.597 | 59.472 | 59.32 | 59.135 | 59.209 | 59.125 | 59.083 | 59.065
| DE to Gretag XYZ.: | 0.019106
Argyll | DE to Gretag spec.: | 0.015709 |  | 30.2755 | 20.074 | 10.7254 | 51.9208 | 47.0807 | 15.7835 | 9.699 | 12.346 | 14.507 | 14.883 | 14.306 | 13.791 | 13.541 | 13.339 | 13.09 | 12.865 | 12.345 | 11.434 | 10.747 | 10.344 | 9.745 | 9.257 | 9.375 | 10.089 | 10.841 | 11.594 | 15.488 | 25.887 | 39.435 | 49.987 | 55.697 | 58.121 | 59.112 | 59.398 | 59.571 | 59.439 | 59.304 | 59.112 | 59.188 | 59.126 | 59.088 | 59.071
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.2446 |  | 30.2619 | 20.0251 | 10.742 | 51.8655 | 47.2679 | 15.6361 | 9.086 | 9.5 | 10.271 | 11.23 | 12.139 | 13.252 | 14.117 | 14.674 | 15.006 | 15.063 | 15.016 | 14.826 | 14.571 | 14.381 | 14.148 | 13.94 | 13.782 | 13.671 | 13.582 | 13.562 | 13.506 | 13.441 | 13.347 | 13.273 | 13.182 | 13.087 | 13.038 | 12.948 | 12.913 | 12.819 | 12.662 | 12.429 | 12.113 | 11.752 | 11.416 | 11.124 | 10.879 | 10.711 | 10.594 | 10.507 | 10.389 | 10.212 | 10.006 | 9.764 | 9.527 | 9.317 | 9.204 | 9.176 | 9.218 | 9.323 | 9.496 | 9.747 | 10.075 | 10.42 | 10.706 | 10.902 | 11.026 | 11.128 | 11.367 | 11.91 | 13.019 | 14.895 | 17.613 | 21.22 | 25.559 | 30.273 | 35.031 | 39.585 | 43.763 | 47.392 | 50.365 | 52.721 | 54.576 | 55.963 | 56.988 | 57.713 | 58.24 | 58.621 | 58.917 | 59.13 | 59.242 | 59.277 | 59.391 | 59.657 | 59.718 | 59.599 | 59.505 | 59.442 | 59.391 | 59.379 | 59.374 | 59.318 | 59.195 | 59.11 | 59.102 | 59.176 | 59.2 | 59.207 | 59.189 | 59.133 | 59.085 | 59.12 | 59.133 | 59.116 | 59.081 | 59.044 | 59.053 | 59.107 | 59.077 | 59.058 | 59.093 | 59.094 | 58.722
Purple
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 8.2658 | 6.3913 | 10.7437 | 30.379 | 20.5567 | -21.4151
Gretag spec. | DE spec Lab: | 0.0041616 |  | 8.2663 | 6.3913 | 10.7438 | 30.3791 | 20.5609 | -21.4154 | 8.458 | 12.078 | 16.398 | 19.225 | 19.715 | 18.656 | 16.646 | 14.304 | 12.004 | 10.043 | 8.614 | 7.515 | 6.607 | 6.026 | 5.685 | 5.378 | 5.18 | 5.183 | 5.278 | 5.299 | 5.156 | 5.174 | 5.688 | 6.987 | 9.128 | 11.59 | 13.858 | 15.862 | 18 | 20.353 | 23.077 | 26.295 | 29.828 | 33.487 | 37.54 | 42.52
| DE to Gretag XYZ.: | 0.039064
Argyll | DE to Gretag spec.: | 0.036791 |  | 8.2801 | 6.4015 | 10.7641 | 30.4035 | 20.5779 | -21.4369 | 8.72 | 12.063 | 16.352 | 19.299 | 19.756 | 18.678 | 16.669 | 14.336 | 12.046 | 10.055 | 8.619 | 7.522 | 6.621 | 6.032 | 5.701 | 5.382 | 5.184 | 5.181 | 5.287 | 5.317 | 5.171 | 5.191 | 5.694 | 6.997 | 9.142 | 11.605 | 13.87 | 15.891 | 18.012 | 20.362 | 23.102 | 26.327 | 29.855 | 33.507 | 37.587 | 42.545
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.17243 |  | 8.2739 | 6.3924 | 10.8051 | 30.3817 | 20.6176 | -21.6033 | 7.732 | 8.191 | 8.965 | 10.213 | 11.652 | 13.227 | 14.922 | 16.433 | 17.788 | 18.775 | 19.432 | 19.821 | 19.936 | 19.911 | 19.74 | 19.356 | 18.81 | 18.209 | 17.466 | 16.707 | 15.928 | 15.148 | 14.39 | 13.615 | 12.846 | 12.077 | 11.384 | 10.72 | 10.089 | 9.538 | 9.068 | 8.647 | 8.236 | 7.862 | 7.519 | 7.194 | 6.871 | 6.575 | 6.341 | 6.171 | 6.027 | 5.901 | 5.798 | 5.718 | 5.618 | 5.492 | 5.383 | 5.299 | 5.224 | 5.169 | 5.145 | 5.146 | 5.163 | 5.203 | 5.252 | 5.295 | 5.333 | 5.351 | 5.329 | 5.278 | 5.22 | 5.158 | 5.121 | 5.12 | 5.145 | 5.237 | 5.395 | 5.619 | 5.941 | 6.366 | 6.911 | 7.558 | 8.296 | 9.112 | 9.972 | 10.824 | 11.648 | 12.441 | 13.209 | 13.927 | 14.594 | 15.229 | 15.906 | 16.623 | 17.34 | 18.031 | 18.745 | 19.518 | 20.333 | 21.199 | 22.124 | 23.072 | 24.087 | 25.163 | 26.314 | 27.523 | 28.714 | 29.891 | 31.038 | 32.216 | 33.455 | 34.807 | 36.186 | 37.636 | 39.172 | 40.82 | 42.584 | 44.438 | 46.326 | 48.296 | 50.225 | 51.875 | 52.939
Yellow Green
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 35.5441 | 44.36 | 8.7852 | 72.4684 | -22.8177 | 57.7306
Gretag spec. | DE spec Lab: | 0.0064648 |  | 35.5454 | 44.3595 | 8.7843 | 72.4681 | -22.8119 | 57.7334 | 5.911 | 6.108 | 6.34 | 6.351 | 6.42 | 6.634 | 6.942 | 7.492 | 8.532 | 10.263 | 13.057 | 17.603 | 25.25 | 36.283 | 47.429 | 53.939 | 55.933 | 55.146 | 52.728 | 49.845 | 46.269 | 42.358 | 38.305 | 35.236 | 33.502 | 32.617 | 32.113 | 31.832 | 32.235 | 33.295 | 34.683 | 35.893 | 36.852 | 37.064 | 36.723 | 36.808
| DE to Gretag XYZ.: | 0.026561
Argyll | DE to Gretag spec.: | 0.032167 |  | 35.5419 | 44.3622 | 8.7923 | 72.4699 | -22.8313 | 57.7078 | 5.841 | 6.062 | 6.313 | 6.423 | 6.489 | 6.611 | 6.921 | 7.511 | 8.533 | 10.282 | 13.073 | 17.635 | 25.276 | 36.316 | 47.472 | 53.956 | 55.944 | 55.15 | 52.709 | 49.813 | 46.263 | 42.342 | 38.3 | 35.235 | 33.501 | 32.613 | 32.118 | 31.832 | 32.248 | 33.315 | 34.694 | 35.917 | 36.854 | 37.069 | 36.742 | 36.835
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.35346 |  | 35.5367 | 44.4036 | 8.7153 | 72.4974 | -22.9675 | 58.0328 | 5.89 | 5.857 | 5.972 | 5.997 | 6.038 | 6.149 | 6.155 | 6.27 | 6.357 | 6.457 | 6.493 | 6.424 | 6.41 | 6.446 | 6.483 | 6.523 | 6.61 | 6.7 | 6.795 | 6.908 | 7.063 | 7.224 | 7.438 | 7.717 | 8.041 | 8.419 | 8.891 | 9.428 | 10.084 | 10.869 | 11.787 | 12.86 | 14.096 | 15.618 | 17.365 | 19.436 | 21.956 | 24.961 | 28.404 | 32.25 | 36.257 | 40.342 | 44.299 | 47.696 | 50.423 | 52.574 | 54.19 | 55.282 | 55.947 | 56.197 | 56.138 | 55.79 | 55.251 | 54.572 | 53.738 | 52.777 | 51.802 | 50.845 | 49.837 | 48.691 | 47.485 | 46.301 | 45.089 | 43.78 | 42.406 | 40.981 | 39.589 | 38.241 | 37.02 | 35.977 | 35.141 | 34.441 | 33.864 | 33.431 | 33.1 | 32.83 | 32.602 | 32.414 | 32.243 | 32.105 | 31.935 | 31.791 | 31.764 | 31.917 | 32.051 | 32.196 | 32.488 | 32.839 | 33.265 | 33.748 | 34.256 | 34.714 | 35.135 | 35.537 | 35.951 | 36.349 | 36.672 | 36.927 | 37.065 | 37.119 | 37.105 | 37.025 | 36.894 | 36.739 | 36.601 | 36.605 | 36.744 | 37.043 | 37.486 | 38.125 | 38.829 | 39.493 | 39.887
Orange Yellow
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 49.3471 | 43.7745 | 6.1933 | 72.0775 | 20.3015 | 67.4848
Gretag spec. | DE spec Lab: | 0.0058338 |  | 49.3496 | 43.7754 | 6.1927 | 72.0781 | 20.3058 | 67.4888 | 5.848 | 5.934 | 6.111 | 6.153 | 6.148 | 6.141 | 6.167 | 6.279 | 6.6 | 7.364 | 8.632 | 10.155 | 11.703 | 13.834 | 18.085 | 25.55 | 35.125 | 44.205 | 50.58 | 55.098 | 58.04 | 60.213 | 61.862 | 63.174 | 64.235 | 65.143 | 66.092 | 66.889 | 67.758 | 68.288 | 68.78 | 69.205 | 69.892 | 70.398 | 70.867 | 71.407
| DE to Gretag XYZ.: | 0.017492
Argyll | DE to Gretag spec.: | 0.02291 |  | 49.3509 | 43.7813 | 6.1981 | 72.0821 | 20.2922 | 67.4707 | 5.893 | 5.944 | 6.088 | 6.143 | 6.107 | 6.105 | 6.187 | 6.304 | 6.625 | 7.359 | 8.629 | 10.166 | 11.713 | 13.85 | 18.088 | 25.569 | 35.151 | 44.221 | 50.586 | 55.099 | 58.024 | 60.205 | 61.855 | 63.175 | 64.239 | 65.153 | 66.096 | 66.891 | 67.74 | 68.283 | 68.786 | 69.221 | 69.895 | 70.425 | 70.907 | 71.418
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.2467 |  | 49.3541 | 43.7312 | 6.1436 | 72.0485 | 20.4457 | 67.6609 | 6.265 | 6.001 | 5.796 | 5.838 | 5.947 | 5.963 | 6.061 | 6.188 | 6.176 | 6.15 | 6.143 | 6.103 | 6.066 | 6.076 | 6.083 | 6.047 | 6.067 | 6.096 | 6.119 | 6.126 | 6.157 | 6.205 | 6.266 | 6.328 | 6.422 | 6.546 | 6.715 | 6.948 | 7.243 | 7.622 | 8.065 | 8.554 | 9.078 | 9.624 | 10.154 | 10.637 | 11.131 | 11.647 | 12.214 | 12.877 | 13.665 | 14.692 | 16.027 | 17.685 | 19.675 | 22.114 | 25.016 | 28.198 | 31.571 | 35.047 | 38.399 | 41.537 | 44.419 | 46.893 | 48.969 | 50.78 | 52.436 | 53.928 | 55.191 | 56.252 | 57.216 | 58.097 | 58.887 | 59.609 | 60.254 | 60.824 | 61.348 | 61.828 | 62.278 | 62.725 | 63.156 | 63.564 | 63.902 | 64.228 | 64.547 | 64.858 | 65.153 | 65.439 | 65.774 | 66.07 | 66.307 | 66.538 | 66.881 | 67.387 | 67.68 | 67.768 | 67.874 | 68.035 | 68.211 | 68.403 | 68.603 | 68.775 | 68.865 | 68.993 | 69.237 | 69.487 | 69.678 | 69.878 | 70.045 | 70.194 | 70.346 | 70.569 | 70.792 | 70.932 | 71.05 | 71.215 | 71.413 | 71.639 | 71.785 | 71.9 | 72.051 | 72.13 | 71.631
Blue
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 7.4254 | 6.0447 | 23.6561 | 29.5251 | 16.4977 | -53.3982
Gretag spec. | DE spec Lab: | 0.011687 |  | 7.4246 | 6.0445 | 23.6502 | 29.5246 | 16.4918 | -53.3881 | 7.174 | 9.547 | 13.31 | 17.718 | 21.303 | 24.921 | 29.534 | 34.477 | 35.616 | 32.147 | 26.013 | 19.04 | 13.197 | 9.194 | 6.795 | 5.519 | 4.879 | 4.473 | 4.201 | 4.045 | 3.966 | 3.912 | 3.889 | 3.88 | 3.911 | 3.968 | 4.042 | 4.126 | 4.229 | 4.299 | 4.336 | 4.385 | 4.515 | 4.775 | 5.134 | 5.681
| DE to Gretag XYZ.: | 0.020818
Argyll | DE to Gretag spec.: | 0.029629 |  | 7.4332 | 6.0513 | 23.6779 | 29.5414 | 16.5014 | -53.4105 | 6.89 | 9.477 | 13.37 | 17.766 | 21.361 | 25.003 | 29.556 | 34.504 | 35.647 | 32.183 | 26.039 | 19.057 | 13.211 | 9.215 | 6.803 | 5.53 | 4.881 | 4.478 | 4.202 | 4.051 | 3.963 | 3.919 | 3.897 | 3.891 | 3.917 | 3.967 | 4.047 | 4.122 | 4.222 | 4.303 | 4.349 | 4.395 | 4.53 | 4.779 | 5.141 | 5.711
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.33031 |  | 7.4488 | 6.0443 | 23.7924 | 29.524 | 16.7253 | -53.6528 | 6.263 | 6.785 | 7.261 | 8.053 | 9.082 | 10.378 | 11.735 | 13.213 | 14.775 | 16.26 | 17.675 | 19.071 | 20.201 | 21.107 | 22.138 | 23.385 | 24.777 | 26.137 | 27.603 | 29.384 | 31.324 | 33.283 | 34.823 | 35.825 | 36.25 | 36.084 | 35.393 | 34.215 | 32.677 | 30.813 | 28.639 | 26.296 | 23.87 | 21.342 | 18.988 | 16.822 | 14.809 | 13.033 | 11.508 | 10.214 | 9.115 | 8.168 | 7.385 | 6.775 | 6.266 | 5.833 | 5.5 | 5.259 | 5.045 | 4.866 | 4.723 | 4.59 | 4.472 | 4.366 | 4.276 | 4.194 | 4.123 | 4.078 | 4.044 | 4.002 | 3.973 | 3.955 | 3.946 | 3.934 | 3.909 | 3.899 | 3.899 | 3.898 | 3.889 | 3.881 | 3.889 | 3.895 | 3.908 | 3.921 | 3.931 | 3.952 | 3.974 | 3.991 | 4.02 | 4.058 | 4.075 | 4.095 | 4.126 | 4.173 | 4.216 | 4.234 | 4.255 | 4.279 | 4.308 | 4.318 | 4.33 | 4.337 | 4.351 | 4.359 | 4.387 | 4.433 | 4.473 | 4.524 | 4.594 | 4.68 | 4.779 | 4.896 | 5.013 | 5.156 | 5.295 | 5.468 | 5.666 | 5.926 | 6.262 | 6.667 | 7.15 | 7.69 | 8.194
Green
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 15.2734 | 23.7217 | 7.659 | 55.8081 | -38.981 | 33.2435
Gretag spec. | DE spec Lab: | 0.0053904 |  | 15.2736 | 23.721 | 7.6581 | 55.8073 | -38.9762 | 33.2458 | 5.555 | 5.978 | 6.231 | 6.354 | 6.37 | 6.49 | 6.781 | 7.246 | 8.027 | 9.386 | 11.386 | 14.127 | 18.06 | 24.178 | 31.836 | 36.081 | 36.048 | 33.311 | 29.147 | 24.853 | 20.665 | 16.751 | 13.176 | 10.566 | 9.094 | 8.37 | 7.956 | 7.615 | 7.434 | 7.445 | 7.653 | 7.976 | 8.425 | 8.805 | 8.887 | 8.806
| DE to Gretag XYZ.: | 0.030965
Argyll | DE to Gretag spec.: | 0.031393 |  | 15.2775 | 23.73 | 7.6557 | 55.8164 | -38.9925 | 33.271 | 5.616 | 5.941 | 6.175 | 6.28 | 6.353 | 6.505 | 6.759 | 7.231 | 8.033 | 9.385 | 11.392 | 14.142 | 18.075 | 24.204 | 31.86 | 36.103 | 36.057 | 33.317 | 29.148 | 24.854 | 20.672 | 16.753 | 13.177 | 10.572 | 9.112 | 8.381 | 7.956 | 7.616 | 7.431 | 7.439 | 7.652 | 7.989 | 8.447 | 8.812 | 8.87 | 8.798
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.30816 |  | 15.2527 | 23.7427 | 7.6024 | 55.8293 | -39.1944 | 33.5034 | 5.409 | 5.544 | 5.644 | 5.709 | 5.808 | 5.995 | 6.111 | 6.218 | 6.239 | 6.23 | 6.283 | 6.306 | 6.343 | 6.391 | 6.385 | 6.431 | 6.517 | 6.574 | 6.647 | 6.734 | 6.854 | 7.03 | 7.192 | 7.389 | 7.666 | 7.94 | 8.311 | 8.721 | 9.233 | 9.831 | 10.511 | 11.269 | 12.084 | 13.023 | 14.036 | 15.142 | 16.416 | 17.861 | 19.583 | 21.603 | 23.963 | 26.652 | 29.485 | 32.03 | 34.031 | 35.474 | 36.38 | 36.751 | 36.709 | 36.301 | 35.571 | 34.579 | 33.4 | 32.079 | 30.653 | 29.127 | 27.608 | 26.182 | 24.798 | 23.378 | 21.977 | 20.636 | 19.348 | 18.054 | 16.755 | 15.464 | 14.247 | 13.117 | 12.079 | 11.196 | 10.471 | 9.887 | 9.411 | 9.049 | 8.76 | 8.542 | 8.36 | 8.203 | 8.065 | 7.949 | 7.82 | 7.686 | 7.59 | 7.525 | 7.468 | 7.407 | 7.392 | 7.396 | 7.426 | 7.47 | 7.55 | 7.64 | 7.737 | 7.845 | 7.97 | 8.128 | 8.293 | 8.456 | 8.598 | 8.734 | 8.819 | 8.891 | 8.907 | 8.893 | 8.864 | 8.808 | 8.782 | 8.809 | 8.856 | 8.962 | 9.116 | 9.265 | 9.412
Red
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 22.6175 | 12.84 | 4.0516 | 42.5211 | 56.1171 | 27.6546
Gretag spec. | DE spec Lab: | 0.0036947 |  | 22.6182 | 12.8402 | 4.0511 | 42.5215 | 56.1185 | 27.658 | 4.994 | 5.053 | 5.162 | 5.126 | 5.098 | 5.003 | 4.992 | 5.015 | 4.975 | 4.831 | 4.715 | 4.614 | 4.55 | 4.469 | 4.454 | 4.475 | 4.549 | 4.674 | 4.93 | 5.55 | 7.016 | 10.868 | 19.158 | 32.468 | 47.923 | 59.936 | 66.581 | 69.608 | 71.229 | 71.988 | 72.547 | 72.948 | 73.493 | 73.757 | 73.948 | 74.216
| DE to Gretag XYZ.: | 0.048006
Argyll | DE to Gretag spec.: | 0.050861 |  | 22.6407 | 12.8583 | 4.0656 | 42.5489 | 56.1025 | 27.6183 | 5.197 | 5.184 | 5.271 | 5.198 | 5.101 | 5.04 | 5.032 | 5.025 | 4.967 | 4.851 | 4.724 | 4.638 | 4.56 | 4.493 | 4.468 | 4.496 | 4.559 | 4.675 | 4.938 | 5.563 | 7.038 | 10.887 | 19.195 | 32.527 | 47.964 | 59.957 | 66.591 | 69.606 | 71.237 | 72.002 | 72.538 | 72.917 | 73.455 | 73.763 | 73.97 | 74.214
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.24061 |  | 22.5791 | 12.7962 | 4.0687 | 42.4546 | 56.2295 | 27.437 | 5.451 | 5.287 | 5.278 | 5.209 | 5.224 | 5.266 | 5.266 | 5.212 | 5.182 | 5.174 | 5.165 | 5.156 | 5.14 | 5.119 | 5.1 | 5.071 | 5.042 | 5.045 | 5.047 | 5.007 | 5.006 | 5.042 | 5.031 | 5.046 | 5.018 | 4.968 | 4.938 | 4.897 | 4.856 | 4.803 | 4.77 | 4.73 | 4.703 | 4.659 | 4.637 | 4.598 | 4.572 | 4.554 | 4.522 | 4.513 | 4.488 | 4.478 | 4.461 | 4.465 | 4.468 | 4.47 | 4.489 | 4.513 | 4.534 | 4.548 | 4.577 | 4.609 | 4.66 | 4.725 | 4.806 | 4.914 | 5.057 | 5.254 | 5.502 | 5.82 | 6.256 | 6.864 | 7.696 | 8.868 | 10.49 | 12.631 | 15.331 | 18.644 | 22.601 | 27.181 | 32.222 | 37.578 | 43.047 | 48.324 | 53.088 | 57.135 | 60.482 | 63.162 | 65.281 | 66.882 | 68.027 | 68.9 | 69.715 | 70.551 | 71.064 | 71.307 | 71.522 | 71.73 | 71.946 | 72.164 | 72.395 | 72.542 | 72.641 | 72.74 | 72.933 | 73.142 | 73.305 | 73.465 | 73.59 | 73.619 | 73.709 | 73.833 | 73.952 | 74.015 | 74.014 | 74.082 | 74.193 | 74.337 | 74.409 | 74.508 | 74.632 | 74.68 | 74.159
Yellow
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 58.7384 | 59.2945 | 7.2139 | 81.4532 | 3.8023 | 79.2487
Gretag spec. | DE spec Lab: | 0.0061041 |  | 58.7411 | 59.295 | 7.2134 | 81.4535 | 3.8078 | 79.2513 | 5.74 | 5.386 | 5.414 | 5.455 | 5.457 | 5.42 | 5.51 | 5.675 | 6.106 | 7.259 | 10.033 | 15.823 | 24.995 | 35.444 | 45.39 | 53.559 | 59.298 | 62.992 | 65.319 | 67.466 | 68.911 | 70.177 | 71.198 | 72.116 | 72.897 | 73.581 | 74.345 | 74.979 | 75.691 | 76.069 | 76.399 | 76.674 | 77.191 | 77.55 | 77.887 | 78.261
| DE to Gretag XYZ.: | 0.011581
Argyll | DE to Gretag spec.: | 0.015694 |  | 58.7401 | 59.2968 | 7.2172 | 81.4545 | 3.8008 | 79.2372 | 5.51 | 5.298 | 5.439 | 5.441 | 5.401 | 5.433 | 5.519 | 5.689 | 6.095 | 7.259 | 10.052 | 15.847 | 25.028 | 35.456 | 45.403 | 53.571 | 59.302 | 62.998 | 65.313 | 67.445 | 68.906 | 70.16 | 71.202 | 72.124 | 72.917 | 73.578 | 74.343 | 74.967 | 75.675 | 76.073 | 76.398 | 76.642 | 77.164 | 77.554 | 77.86 | 78.251
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.30673 |  | 58.7641 | 59.3188 | 7.1478 | 81.4665 | 3.8068 | 79.5437 | 5.909 | 5.34 | 5.396 | 5.312 | 5.264 | 5.324 | 5.359 | 5.438 | 5.441 | 5.393 | 5.37 | 5.36 | 5.415 | 5.454 | 5.418 | 5.423 | 5.442 | 5.418 | 5.455 | 5.529 | 5.573 | 5.607 | 5.678 | 5.761 | 5.858 | 6.038 | 6.271 | 6.583 | 7.052 | 7.682 | 8.547 | 9.72 | 11.202 | 13.157 | 15.521 | 18.279 | 21.486 | 24.956 | 28.516 | 32.06 | 35.482 | 38.865 | 42.233 | 45.365 | 48.246 | 50.974 | 53.529 | 55.792 | 57.77 | 59.459 | 60.866 | 62.035 | 63.074 | 63.965 | 64.72 | 65.408 | 66.13 | 66.863 | 67.507 | 68.024 | 68.488 | 68.945 | 69.409 | 69.836 | 70.235 | 70.58 | 70.9 | 71.208 | 71.516 | 71.832 | 72.148 | 72.452 | 72.71 | 72.95 | 73.162 | 73.388 | 73.61 | 73.834 | 74.1 | 74.368 | 74.52 | 74.668 | 74.995 | 75.472 | 75.699 | 75.729 | 75.816 | 75.904 | 76.021 | 76.174 | 76.318 | 76.413 | 76.461 | 76.511 | 76.689 | 76.883 | 77.034 | 77.191 | 77.291 | 77.4 | 77.521 | 77.68 | 77.829 | 77.909 | 77.989 | 78.099 | 78.258 | 78.411 | 78.543 | 78.681 | 78.765 | 78.711 | 78.271
Magenta
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 30.0301 | 19.7146 | 22.7258 | 51.513 | 47.9178 | -13.7358
Gretag spec. | DE spec Lab: | 0.003785 |  | 30.0311 | 19.7149 | 22.7246 | 51.5133 | 47.9201 | -13.7329 | 11.612 | 17.365 | 26.461 | 33.257 | 35.113 | 34.521 | 32.681 | 30.046 | 27.17 | 24.359 | 21.59 | 18.913 | 16.834 | 15.015 | 12.812 | 10.966 | 10.427 | 10.712 | 10.739 | 11.233 | 13.695 | 18.648 | 26.004 | 36.372 | 48.978 | 60.496 | 68.734 | 73.763 | 77.029 | 78.734 | 79.568 | 79.994 | 80.633 | 81.128 | 81.673 | 82.427
| DE to Gretag XYZ.: | 0.031726
Argyll | DE to Gretag spec.: | 0.033489 |  | 30.0549 | 19.7326 | 22.7568 | 51.5335 | 47.9227 | -13.7595 | 11.655 | 17.376 | 26.438 | 33.255 | 35.097 | 34.585 | 32.713 | 30.098 | 27.227 | 24.391 | 21.613 | 18.954 | 16.857 | 15.035 | 12.829 | 10.965 | 10.438 | 10.721 | 10.758 | 11.245 | 13.707 | 18.667 | 26.034 | 36.4 | 49.008 | 60.523 | 68.77 | 73.806 | 77.046 | 78.732 | 79.598 | 80.033 | 80.69 | 81.153 | 81.673 | 82.436
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.28003 |  | 30.0135 | 19.6757 | 22.8252 | 51.4686 | 48.047 | -14.0019 | 10.27 | 11.09 | 12.285 | 14.027 | 16.518 | 19.572 | 22.941 | 26.486 | 29.613 | 31.975 | 33.513 | 34.401 | 34.996 | 35.31 | 35.336 | 35.162 | 34.787 | 34.231 | 33.513 | 32.781 | 31.97 | 31.096 | 30.159 | 29.216 | 28.236 | 27.311 | 26.395 | 25.449 | 24.496 | 23.605 | 22.656 | 21.7 | 20.743 | 19.77 | 18.91 | 18.139 | 17.426 | 16.81 | 16.257 | 15.692 | 15.106 | 14.418 | 13.65 | 12.852 | 12.094 | 11.397 | 10.868 | 10.546 | 10.38 | 10.361 | 10.466 | 10.616 | 10.754 | 10.834 | 10.809 | 10.734 | 10.693 | 10.793 | 11.078 | 11.591 | 12.402 | 13.512 | 14.897 | 16.546 | 18.464 | 20.622 | 23.008 | 25.684 | 28.756 | 32.268 | 36.131 | 40.302 | 44.687 | 49.132 | 53.398 | 57.338 | 60.885 | 64.023 | 66.777 | 69.08 | 70.942 | 72.52 | 73.996 | 75.458 | 76.494 | 77.195 | 77.777 | 78.295 | 78.738 | 79.134 | 79.422 | 79.627 | 79.711 | 79.834 | 80.024 | 80.273 | 80.488 | 80.699 | 80.821 | 80.939 | 81.095 | 81.276 | 81.506 | 81.706 | 81.893 | 82.12 | 82.451 | 82.767 | 83.054 | 83.334 | 83.618 | 83.747 | 83.358
Cyan
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 13.4126 | 19.2483 | 30.1692 | 50.9764 | -29.6205 | -27.5503
Gretag spec. | DE spec Lab: | 0.007642 |  | 13.4119 | 19.2478 | 30.1641 | 50.9759 | -29.6228 | -27.543 | 9.935 | 13.795 | 19.325 | 24.107 | 26.491 | 28.646 | 31.479 | 34.994 | 38.834 | 42.68 | 44.795 | 44.917 | 43.082 | 39.416 | 34.429 | 28.876 | 23.778 | 18.957 | 14.726 | 11.709 | 9.836 | 8.737 | 7.965 | 7.451 | 7.218 | 7.142 | 7.159 | 7.271 | 7.488 | 7.643 | 7.581 | 7.399 | 7.174 | 7.065 | 7.262 | 7.865
| DE to Gretag XYZ.: | 0.02699
Argyll | DE to Gretag spec.: | 0.022472 |  | 13.4218 | 19.2628 | 30.1772 | 50.9933 | -29.6338 | -27.5339 | 9.795 | 13.874 | 19.407 | 24.091 | 26.465 | 28.641 | 31.485 | 35.025 | 38.851 | 42.706 | 44.82 | 44.945 | 43.088 | 39.451 | 34.455 | 28.897 | 23.794 | 18.97 | 14.746 | 11.724 | 9.84 | 8.748 | 7.975 | 7.455 | 7.224 | 7.148 | 7.165 | 7.281 | 7.499 | 7.642 | 7.597 | 7.403 | 7.181 | 7.058 | 7.266 | 7.878
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.25068 |  | 13.3938 | 19.2742 | 30.1414 | 51.0064 | -29.8712 | -27.4546 | 8.776 | 9.207 | 10.289 | 11.588 | 13.153 | 15.183 | 17.36 | 19.417 | 21.277 | 22.914 | 24.272 | 25.202 | 25.799 | 26.329 | 27.029 | 27.853 | 28.518 | 29.242 | 30.272 | 31.435 | 32.655 | 33.829 | 34.989 | 36.151 | 37.332 | 38.634 | 40.058 | 41.447 | 42.705 | 43.705 | 44.453 | 44.984 | 45.233 | 45.258 | 45.113 | 44.723 | 44.044 | 43.216 | 42.196 | 40.98 | 39.595 | 38.048 | 36.355 | 34.648 | 32.82 | 30.884 | 29.021 | 27.281 | 25.578 | 23.856 | 22.194 | 20.539 | 18.917 | 17.389 | 15.957 | 14.619 | 13.435 | 12.439 | 11.618 | 10.894 | 10.264 | 9.777 | 9.379 | 9.036 | 8.732 | 8.465 | 8.201 | 7.963 | 7.746 | 7.566 | 7.423 | 7.337 | 7.264 | 7.216 | 7.187 | 7.168 | 7.15 | 7.149 | 7.144 | 7.153 | 7.19 | 7.225 | 7.276 | 7.362 | 7.447 | 7.522 | 7.594 | 7.631 | 7.654 | 7.654 | 7.642 | 7.605 | 7.552 | 7.47 | 7.404 | 7.347 | 7.255 | 7.158 | 7.087 | 7.032 | 7.02 | 7.074 | 7.149 | 7.262 | 7.415 | 7.586 | 7.829 | 8.163 | 8.551 | 9.079 | 9.693 | 10.388 | 11.101
White
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 85.7926 | 89.3218 | 70.2957 | 95.7148 | -0.61862 | 2.9967
Gretag spec. | DE spec Lab: | 0.0081723 |  | 85.7949 | 89.3219 | 70.2879 | 95.7148 | -0.61442 | 3.0037 | 14.809 | 22.568 | 40.024 | 64.348 | 79.036 | 83.188 | 84.436 | 85.453 | 86.408 | 87.003 | 87.443 | 87.938 | 88.372 | 88.613 | 88.935 | 89.13 | 89.417 | 89.548 | 89.377 | 89.72 | 89.621 | 89.643 | 89.689 | 89.735 | 89.769 | 89.793 | 89.965 | 90.024 | 90.239 | 90.164 | 90.083 | 89.937 | 90.153 | 90.206 | 90.211 | 90.325
| DE to Gretag XYZ.: | 0.022168
Argyll | DE to Gretag spec.: | 0.015894 |  | 85.77 | 89.2935 | 70.2546 | 95.7029 | -0.60988 | 3.0133 | 14.815 | 22.631 | 39.96 | 64.212 | 78.941 | 83.105 | 84.427 | 85.445 | 86.359 | 86.957 | 87.387 | 87.91 | 88.328 | 88.589 | 88.928 | 89.105 | 89.375 | 89.499 | 89.339 | 89.684 | 89.599 | 89.641 | 89.664 | 89.721 | 89.775 | 89.749 | 89.921 | 90.006 | 90.226 | 90.138 | 90.057 | 89.929 | 90.115 | 90.173 | 90.175 | 90.3
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.027963 |  | 85.7492 | 89.2638 | 70.2537 | 95.6906 | -0.59545 | 2.9928 | 13.647 | 13.977 | 15.32 | 17.582 | 20.971 | 25.642 | 31.6 | 38.879 | 47.31 | 56.133 | 64.511 | 71.412 | 76.27 | 79.427 | 81.32 | 82.437 | 83.201 | 83.705 | 84.057 | 84.383 | 84.746 | 85.107 | 85.477 | 85.792 | 86.091 | 86.329 | 86.57 | 86.748 | 86.911 | 87.062 | 87.198 | 87.363 | 87.549 | 87.704 | 87.899 | 88.052 | 88.178 | 88.268 | 88.4 | 88.509 | 88.576 | 88.681 | 88.772 | 88.908 | 88.96 | 88.936 | 89.009 | 89.16 | 89.304 | 89.392 | 89.432 | 89.443 | 89.426 | 89.4 | 89.356 | 89.319 | 89.404 | 89.541 | 89.65 | 89.649 | 89.58 | 89.571 | 89.61 | 89.646 | 89.652 | 89.636 | 89.61 | 89.617 | 89.603 | 89.636 | 89.68 | 89.729 | 89.756 | 89.753 | 89.742 | 89.736 | 89.744 | 89.772 | 89.839 | 89.881 | 89.854 | 89.82 | 90.002 | 90.368 | 90.412 | 90.237 | 90.119 | 90.045 | 90.026 | 90.069 | 90.076 | 90.06 | 89.988 | 89.933 | 89.983 | 90.019 | 90.052 | 90.111 | 90.131 | 90.121 | 90.133 | 90.153 | 90.213 | 90.206 | 90.192 | 90.215 | 90.292 | 90.373 | 90.414 | 90.386 | 90.454 | 90.44 | 90.083
Neutral 8
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 56.3747 | 58.5518 | 47.8431 | 81.0446 | -0.20003 | 0.52851
Gretag spec. | DE spec Lab: | 0.0070775 |  | 56.3762 | 58.5518 | 47.8379 | 81.0446 | -0.19643 | 0.5346 | 13.954 | 21.246 | 35.511 | 50.422 | 56.381 | 57.61 | 57.957 | 58.247 | 58.504 | 58.508 | 58.485 | 58.553 | 58.604 | 58.534 | 58.495 | 58.424 | 58.525 | 58.543 | 58.406 | 58.684 | 58.698 | 58.734 | 58.717 | 58.677 | 58.569 | 58.411 | 58.35 | 58.235 | 58.216 | 58.028 | 57.822 | 57.568 | 57.548 | 57.455 | 57.375 | 57.293
| DE to Gretag XYZ.: | 0.013491
Argyll | DE to Gretag spec.: | 0.019339 |  | 56.3869 | 58.5646 | 47.8635 | 81.0517 | -0.20033 | 0.51702 | 13.967 | 21.295 | 35.548 | 50.561 | 56.432 | 57.646 | 58.011 | 58.3 | 58.502 | 58.529 | 58.509 | 58.561 | 58.614 | 58.541 | 58.504 | 58.435 | 58.543 | 58.568 | 58.429 | 58.704 | 58.716 | 58.75 | 58.707 | 58.666 | 58.57 | 58.411 | 58.358 | 58.24 | 58.204 | 58.014 | 57.836 | 57.584 | 57.584 | 57.468 | 57.391 | 57.325
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.030089 |  | 56.3954 | 58.5676 | 47.8889 | 81.0533 | -0.1864 | 0.49041 | 12.806 | 13.09 | 14.519 | 16.688 | 19.906 | 23.991 | 29.205 | 35.089 | 41.18 | 46.792 | 51.228 | 54.147 | 55.807 | 56.716 | 57.194 | 57.449 | 57.659 | 57.871 | 57.918 | 57.989 | 58.092 | 58.249 | 58.378 | 58.454 | 58.506 | 58.538 | 58.532 | 58.53 | 58.547 | 58.528 | 58.515 | 58.543 | 58.521 | 58.529 | 58.564 | 58.59 | 58.592 | 58.604 | 58.595 | 58.593 | 58.579 | 58.526 | 58.498 | 58.543 | 58.481 | 58.401 | 58.403 | 58.48 | 58.546 | 58.586 | 58.583 | 58.557 | 58.539 | 58.507 | 58.466 | 58.431 | 58.473 | 58.589 | 58.698 | 58.709 | 58.69 | 58.728 | 58.767 | 58.784 | 58.778 | 58.758 | 58.731 | 58.709 | 58.672 | 58.678 | 58.678 | 58.668 | 58.631 | 58.576 | 58.523 | 58.472 | 58.417 | 58.385 | 58.374 | 58.353 | 58.29 | 58.2 | 58.237 | 58.384 | 58.372 | 58.217 | 58.109 | 58.02 | 57.98 | 57.932 | 57.899 | 57.835 | 57.721 | 57.617 | 57.589 | 57.597 | 57.601 | 57.596 | 57.529 | 57.457 | 57.432 | 57.438 | 57.43 | 57.434 | 57.378 | 57.338 | 57.33 | 57.35 | 57.305 | 57.283 | 57.327 | 57.309 | 57.047
Neutral 6.5
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 34.2366 | 35.4994 | 28.9469 | 66.1356 | 0.028027 | 0.54469
Gretag spec. | DE spec Lab: | 0.0061445 |  | 34.2375 | 35.4994 | 28.9436 | 66.1357 | 0.030997 | 0.55007 | 12.416 | 17.942 | 26.257 | 32.441 | 34.236 | 34.716 | 35.023 | 35.257 | 35.378 | 35.345 | 35.263 | 35.281 | 35.323 | 35.305 | 35.318 | 35.314 | 35.41 | 35.435 | 35.409 | 35.663 | 35.762 | 35.818 | 35.797 | 35.738 | 35.601 | 35.441 | 35.356 | 35.19 | 35.101 | 34.894 | 34.697 | 34.475 | 34.416 | 34.248 | 34.15 | 34.033
| DE to Gretag XYZ.: | 0.010801
Argyll | DE to Gretag spec.: | 0.016487 |  | 34.2305 | 35.4955 | 28.9476 | 66.1326 | 0.019877 | 0.53829 | 12.219 | 17.951 | 26.197 | 32.365 | 34.211 | 34.73 | 35.016 | 35.268 | 35.389 | 35.347 | 35.281 | 35.292 | 35.341 | 35.312 | 35.323 | 35.316 | 35.401 | 35.438 | 35.408 | 35.649 | 35.743 | 35.811 | 35.788 | 35.722 | 35.594 | 35.434 | 35.337 | 35.186 | 35.1 | 34.893 | 34.707 | 34.483 | 34.397 | 34.259 | 34.161 | 34.032
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.014036 |  | 34.2429 | 35.5071 | 28.9633 | 66.1416 | 0.023881 | 0.52829 | 10.736 | 11.71 | 12.892 | 14.792 | 17.122 | 19.956 | 23.189 | 26.304 | 29.035 | 31.192 | 32.608 | 33.516 | 34.044 | 34.272 | 34.466 | 34.595 | 34.713 | 34.856 | 34.887 | 34.998 | 35.102 | 35.207 | 35.317 | 35.41 | 35.411 | 35.396 | 35.443 | 35.37 | 35.348 | 35.321 | 35.33 | 35.335 | 35.305 | 35.286 | 35.304 | 35.325 | 35.32 | 35.337 | 35.358 | 35.367 | 35.334 | 35.313 | 35.322 | 35.35 | 35.333 | 35.275 | 35.292 | 35.361 | 35.418 | 35.451 | 35.441 | 35.421 | 35.414 | 35.422 | 35.423 | 35.409 | 35.462 | 35.579 | 35.664 | 35.699 | 35.723 | 35.776 | 35.818 | 35.833 | 35.837 | 35.841 | 35.82 | 35.803 | 35.768 | 35.741 | 35.739 | 35.707 | 35.653 | 35.596 | 35.558 | 35.508 | 35.441 | 35.391 | 35.368 | 35.34 | 35.272 | 35.199 | 35.171 | 35.253 | 35.237 | 35.097 | 35.009 | 34.93 | 34.843 | 34.828 | 34.786 | 34.717 | 34.615 | 34.537 | 34.488 | 34.475 | 34.448 | 34.42 | 34.362 | 34.281 | 34.225 | 34.226 | 34.214 | 34.161 | 34.116 | 34.088 | 34.058 | 34.017 | 33.968 | 33.9 | 33.886 | 33.845 | 33.674
Neutral 5
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 18.5614 | 19.2728 | 16.0738 | 51.0049 | -0.11105 | -0.42399
Gretag spec. | DE spec Lab: | 0.0049883 |  | 18.5619 | 19.2728 | 16.072 | 51.0049 | -0.10868 | -0.4196 | 10.034 | 13.202 | 16.805 | 18.803 | 19.26 | 19.422 | 19.497 | 19.616 | 19.602 | 19.529 | 19.449 | 19.402 | 19.386 | 19.328 | 19.283 | 19.248 | 19.275 | 19.252 | 19.21 | 19.331 | 19.358 | 19.36 | 19.319 | 19.269 | 19.171 | 19.041 | 18.94 | 18.822 | 18.743 | 18.594 | 18.467 | 18.325 | 18.252 | 18.155 | 18.07 | 17.989
| DE to Gretag XYZ.: | 0.0298
Argyll | DE to Gretag spec.: | 0.027693 |  | 18.5739 | 19.2878 | 16.0771 | 51.0223 | -0.1211 | -0.40198 | 9.995 | 13.253 | 16.786 | 18.79 | 19.251 | 19.422 | 19.519 | 19.602 | 19.605 | 19.54 | 19.47 | 19.422 | 19.399 | 19.342 | 19.302 | 19.27 | 19.291 | 19.269 | 19.226 | 19.333 | 19.373 | 19.379 | 19.341 | 19.282 | 19.17 | 19.05 | 18.959 | 18.841 | 18.755 | 18.616 | 18.479 | 18.33 | 18.277 | 18.167 | 18.098 | 17.998
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.029029 |  | 18.5726 | 19.2832 | 16.0831 | 51.017 | -0.10496 | -0.42552 | 9.189 | 9.646 | 10.511 | 11.555 | 12.913 | 14.322 | 15.7 | 16.916 | 17.858 | 18.529 | 18.921 | 19.092 | 19.222 | 19.314 | 19.327 | 19.353 | 19.418 | 19.464 | 19.476 | 19.513 | 19.547 | 19.589 | 19.644 | 19.658 | 19.639 | 19.617 | 19.597 | 19.563 | 19.535 | 19.519 | 19.485 | 19.472 | 19.441 | 19.407 | 19.427 | 19.402 | 19.405 | 19.394 | 19.379 | 19.372 | 19.345 | 19.324 | 19.303 | 19.315 | 19.274 | 19.239 | 19.239 | 19.275 | 19.304 | 19.3 | 19.291 | 19.262 | 19.247 | 19.246 | 19.227 | 19.212 | 19.237 | 19.291 | 19.333 | 19.344 | 19.342 | 19.37 | 19.397 | 19.394 | 19.382 | 19.386 | 19.364 | 19.336 | 19.304 | 19.293 | 19.279 | 19.256 | 19.206 | 19.167 | 19.131 | 19.086 | 19.058 | 19.014 | 18.974 | 18.953 | 18.897 | 18.858 | 18.832 | 18.863 | 18.84 | 18.76 | 18.697 | 18.642 | 18.607 | 18.576 | 18.53 | 18.49 | 18.429 | 18.356 | 18.331 | 18.324 | 18.298 | 18.285 | 18.238 | 18.18 | 18.155 | 18.144 | 18.127 | 18.104 | 18.07 | 18.014 | 17.994 | 17.978 | 17.937 | 17.908 | 17.925 | 17.909 | 17.803
Neutral 3.5
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 8.2408 | 8.5484 | 6.9823 | 35.0998 | -0.013786 | 0.28945
Gretag spec. | DE spec Lab: | 0.0038948 |  | 8.241 | 8.5484 | 6.9815 | 35.0998 | -0.011994 | 0.29291 | 6.477 | 7.256 | 7.887 | 8.16 | 8.292 | 8.375 | 8.493 | 8.521 | 8.539 | 8.46 | 8.417 | 8.409 | 8.449 | 8.469 | 8.5 | 8.518 | 8.544 | 8.563 | 8.548 | 8.608 | 8.635 | 8.635 | 8.621 | 8.588 | 8.537 | 8.487 | 8.453 | 8.392 | 8.355 | 8.311 | 8.259 | 8.194 | 8.178 | 8.139 | 8.106 | 8.085
| DE to Gretag XYZ.: | 0.02251
Argyll | DE to Gretag spec.: | 0.024228 |  | 8.2456 | 8.5554 | 6.9881 | 35.1138 | -0.031388 | 0.28912 | 6.409 | 7.204 | 7.934 | 8.225 | 8.265 | 8.376 | 8.48 | 8.555 | 8.535 | 8.475 | 8.427 | 8.432 | 8.458 | 8.482 | 8.511 | 8.526 | 8.558 | 8.564 | 8.557 | 8.614 | 8.634 | 8.646 | 8.627 | 8.589 | 8.542 | 8.489 | 8.447 | 8.398 | 8.365 | 8.304 | 8.253 | 8.191 | 8.173 | 8.141 | 8.111 | 8.075
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.01427 |  | 8.2509 | 8.5597 | 6.993 | 35.1224 | -0.021439 | 0.28362 | 6.357 | 6.319 | 6.556 | 6.873 | 7.172 | 7.552 | 7.747 | 7.922 | 8.026 | 8.104 | 8.223 | 8.241 | 8.23 | 8.295 | 8.318 | 8.327 | 8.386 | 8.411 | 8.438 | 8.466 | 8.509 | 8.553 | 8.568 | 8.585 | 8.56 | 8.546 | 8.52 | 8.506 | 8.485 | 8.464 | 8.449 | 8.434 | 8.427 | 8.427 | 8.447 | 8.441 | 8.444 | 8.454 | 8.472 | 8.484 | 8.495 | 8.5 | 8.498 | 8.513 | 8.516 | 8.514 | 8.518 | 8.539 | 8.551 | 8.57 | 8.57 | 8.569 | 8.564 | 8.566 | 8.565 | 8.554 | 8.575 | 8.602 | 8.621 | 8.633 | 8.625 | 8.633 | 8.659 | 8.661 | 8.65 | 8.645 | 8.638 | 8.623 | 8.621 | 8.606 | 8.592 | 8.588 | 8.571 | 8.547 | 8.541 | 8.526 | 8.496 | 8.473 | 8.463 | 8.458 | 8.429 | 8.418 | 8.406 | 8.414 | 8.389 | 8.354 | 8.333 | 8.312 | 8.291 | 8.289 | 8.288 | 8.26 | 8.248 | 8.223 | 8.208 | 8.183 | 8.179 | 8.19 | 8.156 | 8.137 | 8.135 | 8.132 | 8.132 | 8.131 | 8.109 | 8.094 | 8.097 | 8.07 | 8.056 | 8.062 | 8.053 | 8.021 | 7.994
Black
|  | DE's | Reading | X | Y | Z | L* | a* | b* | 380 | 390 | 400 | 410 | 420 | 430 | 440 | 450 | 460 | 470 | 480 | 490 | 500 | 510 | 520 | 530 | 540 | 550 | 560 | 570 | 580 | 590 | 600 | 610 | 620 | 630 | 640 | 650 | 660 | 670 | 680 | 690 | 700 | 710 | 720 | 730
Gretag XYZ |  |  |  | 3.0298 | 3.1199 | 2.5618 | 20.5179 | 0.37595 | 0.096673
Gretag spec. | DE spec Lab: | 0.0027344 |  | 3.0299 | 3.1199 | 2.5615 | 20.5179 | 0.37727 | 0.099066 | 3.132 | 3.166 | 3.12 | 3.108 | 3.143 | 3.098 | 3.1 | 3.113 | 3.108 | 3.087 | 3.105 | 3.104 | 3.101 | 3.099 | 3.096 | 3.09 | 3.104 | 3.1 | 3.1 | 3.116 | 3.123 | 3.138 | 3.143 | 3.155 | 3.169 | 3.182 | 3.21 | 3.232 | 3.264 | 3.274 | 3.289 | 3.312 | 3.342 | 3.377 | 3.391 | 3.41
| DE to Gretag XYZ.: | 0.043154
Argyll | DE to Gretag spec.: | 0.043664 |  | 3.0388 | 3.1279 | 2.5708 | 20.5488 | 0.39831 | 0.076473 | 3 | 3.115 | 3.174 | 3.147 | 3.148 | 3.137 | 3.128 | 3.115 | 3.107 | 3.103 | 3.098 | 3.099 | 3.111 | 3.116 | 3.106 | 3.099 | 3.107 | 3.106 | 3.106 | 3.122 | 3.131 | 3.144 | 3.158 | 3.17 | 3.179 | 3.191 | 3.213 | 3.23 | 3.263 | 3.28 | 3.296 | 3.32 | 3.354 | 3.382 | 3.403 | 3.437
|  |  | Reading | X | Y | Z | L* | a* | b* | 376.6667 | 380 | 383.3333 | 386.6667 | 390 | 393.3333 | 396.6667 | 400 | 403.3333 | 406.6667 | 410 | 413.3333 | 416.6667 | 420 | 423.3333 | 426.6667 | 430 | 433.3333 | 436.6667 | 440 | 443.3333 | 446.6667 | 450 | 453.3333 | 456.6667 | 460 | 463.3333 | 466.6667 | 470 | 473.3333 | 476.6667 | 480 | 483.3333 | 486.6667 | 490 | 493.3333 | 496.6667 | 500 | 503.3333 | 506.6667 | 510 | 513.3333 | 516.6667 | 520 | 523.3333 | 526.6667 | 530 | 533.3333 | 536.6667 | 540 | 543.3333 | 546.6667 | 550 | 553.3333 | 556.6667 | 560 | 563.3333 | 566.6667 | 570 | 573.3333 | 576.6667 | 580 | 583.3333 | 586.6667 | 590 | 593.3333 | 596.6667 | 600 | 603.3333 | 606.6667 | 610 | 613.3333 | 616.6667 | 620 | 623.3333 | 626.6667 | 630 | 633.3333 | 636.6667 | 640 | 643.3333 | 646.6667 | 650 | 653.3333 | 656.6667 | 660 | 663.3333 | 666.6667 | 670 | 673.3333 | 676.6667 | 680 | 683.3333 | 686.6667 | 690 | 693.3333 | 696.6667 | 700 | 703.3333 | 706.6667 | 710 | 713.3333 | 716.6667 | 720 | 723.3333 | 726.6667 | 730 | 733.3333 | 736.6667 | 740 | 743.3333 | 746.6667 | 750
Argyll HiRes | DE to Argyll Norm: | 0.064678 |  | 3.042 | 3.1299 | 2.58 | 20.557 | 0.418 | 0.015405 | 2.89 | 3.081 | 3.047 | 3.105 | 3.11 | 3.08 | 3.173 | 3.214 | 3.234 | 3.27 | 3.236 | 3.177 | 3.191 | 3.208 | 3.174 | 3.15 | 3.155 | 3.169 | 3.136 | 3.133 | 3.14 | 3.115 | 3.124 | 3.125 | 3.12 | 3.132 | 3.11 | 3.099 | 3.11 | 3.107 | 3.101 | 3.1 | 3.108 | 3.105 | 3.108 | 3.113 | 3.105 | 3.125 | 3.114 | 3.118 | 3.125 | 3.11 | 3.108 | 3.11 | 3.111 | 3.099 | 3.097 | 3.106 | 3.101 | 3.105 | 3.109 | 3.111 | 3.107 | 3.113 | 3.107 | 3.098 | 3.108 | 3.119 | 3.126 | 3.124 | 3.128 | 3.139 | 3.145 | 3.143 | 3.145 | 3.147 | 3.159 | 3.168 | 3.16 | 3.161 | 3.169 | 3.176 | 3.174 | 3.177 | 3.185 | 3.19 | 3.189 | 3.205 | 3.215 | 3.216 | 3.216 | 3.222 | 3.237 | 3.247 | 3.256 | 3.262 | 3.276 | 3.267 | 3.275 | 3.3 | 3.3 | 3.292 | 3.297 | 3.314 | 3.308 | 3.335 | 3.348 | 3.341 | 3.355 | 3.359 | 3.385 | 3.4 | 3.418 | 3.427 | 3.422 | 3.427 | 3.439 | 3.437 | 3.463 | 3.453 | 3.459 | 3.477 | 3.439